## Cal
| QC-based Nested Approach for Estimating Analytical Measurement Uncertainty |  |  |  |  | Page 1 | QC-based Nested Approach | [6 empty] | Copper in Wastewater by ICP | [6 empty] | Page 2 | QC-based Nested Approach |  |  |  |  | Copper in Wastewater by ICP |  |  |  | Page 3
| What are the analyte, matrix, and technology? |  |  | Copper in Wastewater by ICP
| Enter 20 replicate results for the following quality control samples as relative deviation (%): | [6 empty] | Components of Analytical Uncertainty
| ICS - Instrument calibration standard | [6 empty] | IME - Intrinsic instrumental measurement effects
| ICV - Second source calibration verification standard | [6 empty] | SPE - Spike preparation effects
| LCS - Laboratory control sample | [6 empty] | PME - Preparation method effects
| MIS - Matrix interference sample (matrix spike, organic surrogate, radiochemical tracer) | [6 empty] | MIE - Matrix interference effects
| FDS - Field-split duplicate sample | [6 empty] | SCE- Sample collection effects
| CLS - Co-located duplicate sample | [6 empty] | SLE - Sample location effects
| ICS | ICV | LCS | MIS | FDS | CLS |  | Component Percent Standard Uncertainty | [6 empty] | Component Percent Recovery |  |  | Component Systematic Error
| 1.1 | 0.5 | 4 | 12 | 0 | 0 |  | IME | ~ | 0.84195 | % relative standard deviation |  |  |  | IME | ~ | 101.4795 | IME | ~ | 1.4795 | percent
| 0.76 | 0.1 | 0.5 | 1.4 | 0 | 0
| 0.36 | 1 | 1.5 | 8 | 0 | 0 |  | SPE | ~ | 0.096465 | % relative standard deviation |  |  |  | SPE | ~ | 99.6605 | SPE | ~ | -0.33948 | percent
| 2.04 | 1.2 | 1.7 | 3.7 | 0 | 0
| 0.99 | 0.2 | 0.1 | 12 | 0 | 0 |  | PME | ~ | 7.1307 | % relative standard deviation |  |  |  | PME | ~ | 104.2517 | PME | ~ | 4.2517 | percent
| 1.21 | 0.4 | 2.2 | 0.4 | 0 | 0
| 1.74 | 1.2 | 0.4 | 3.6 | 0 | 0 |  | MIE | ~ | 8.5246 | % relative standard deviation |  |  |  | MIE | ~ | 99.3029 | MIE | ~ | -0.69711 | percent
| 3.73 | 0.9 | 0.3 | 0.1 | 0 | 0 | [14 empty] | 2.093
| 1.06 | 0.1 | 0.5 | 2.7 | 0 | 0 |  | SCE | ~ | 0 | % relative standard deviation | [9 empty] | 2.093 |  | What is the analytical measurement result? |  |  |  |  | 10
| 3.1 | 1.3 | 15 | 17 | 0 | 0 | [14 empty] | 2.093
| 2 | 0.9 | 20 | 30 | 0 | 0 |  | SLE | ~ | 0 | % relative standard deviation | [9 empty] | WRONG CL |  | What are the analytical measurement units? |  |  |  |  | mg/L
| 0.7 | 1 | 0.4 | 3.7 | 0 | 0 | [14 empty] | WRONG CL
| 0.4 | 2 | 4 | 1.5 | 0 | 0 |  | What is the Confidence Level (CL)? Enter ONLY one of these percentages: 80, 90, 95, 99 | [10 empty] | 95 | %
| 0.9 | 0.2 | 0.6 | 5 | 0 | 0 |  | Your specified t-value is |  |  |  | 2.093 | for a Two-Tailed Normal Distribution Confidence Interval
| 1.4 | 1 | 1.5 | 1.4 | 0 | 0
| 1.9 | 1.4 | 5 | 20 | 0 | 0 |  | Relative Analytical Measurement Uncertainty for routine field samples | [14 empty] | If the sample measurement is | 10 | mg/L |  | ,
| 2 | 1.5 | 24 | 3.5 | 0 | 0 |  | (Only the IME, PME, and MIE are combined for the analytical measurement uncertainty) | [14 empty] | then the uncertainty interval is | 7.6672 | - | 12.3328 | mg/L | at the | 95 | % Confidence Level (Expanded Uncertainty)
| 1.5 | 1.7 | 3 | 5 | 0 | 0 |  |  |  | 23.3276 | % relative uncertainty
| 1.6 | 3 | 13 | -24 | 0 | 0
| 1.1 | 3.1 | 11 | -13 | 0 | 0 | [16 empty] | For the above result, if the systematic measurement error (bias) is corrected, and
Std. Dev. | 0.84195 | 0.84746 | 7.1808 | 11.146 | 0 | 0 |  | Relative Systematic Error associated with the measurement of routine field samples | [14 empty] | the corrected measurement is | 9.5187 | mg/L |  | ,
Bias | 1.4795 | 1.135 | 5.435 | 4.7 |  |  |  | (Only the IME, PME, and MIE biases are combined for the analytical measurement systematic error) | [14 empty] | then the uncertainty interval is | 7.2982 | - | 11.7392 | mg/L | at the | 95 | % Confidence Level (Expanded Uncertainty)
Recovery | 101.4795 | 101.135 | 105.435 | 104.7 | [5 empty] | 5.0566 | % relative systematic error

## Check
CHECK FOR NORMALITY
CHECK WITH DOCUMENTED DATA
1 | 942 | 941 | 0.4734 | 445.4694
3.1 | 637 | 633.9 | 0.3211 | 203.5453
8.7 | 578 | 569.3 | 0.2565 | 146.0254
10 | 331 | 321 | 0.2085 | 66.9285
14 | 262 | 248 | 0.1686 | 41.8128
19 | 151 | 132 | 0.1334 | 17.6088
21.4 | 85.6 | 64.2 | 0.1013 | 6.5035
27 | 81.5 | 54.5 | 0.0711 | 3.8749
39 | 64.4 | 25.4 | 0.0422 | 1.0719
56 | 58.8 | 2.8 | 0.014 | 0.0392
58.8 | 56 | -2.8 |  | 932.8797
64.4 | 39 | -25.4
81.5 | 27 | -54.5 | W | 0.67904
85.6 | 21.4 | -64.2
151 | 19 | -132
262 | 14 | -248
331 | 10 | -321
578 | 8.7 | -569.3
637 | 3.1 | -633.9
942 | 1 | -941
259.7175

## Shapiro-Wilk Test for Normality
INSTRUMENT CALIBRATION STANDARD | [5 empty] | 1.2 | 0.024649 | 0.13 | 2.06 | 0.5739 | -1.1076
0.36 | 3.73 | 3.37 | 0.4734 | 1.5954 |  | 0.13 | 0.83357 | 0.45 | 1.69 | 0.3291 | -0.40808
0.4 | 3.1 | 2.7 | 0.3211 | 0.86697 |  | 1.69 | 0.41861 | 0.6 | 1.37 | 0.2141 | -0.16486
0.7 | 2.04 | 1.34 | 0.2565 | 0.34371 |  | 1.05 | 4.9e-05 | 0.76 | 1.2 | 0.1224 | -0.053856
0.76 | 2 | 1.24 | 0.2085 | 0.25854 |  | 1.12 | 0.005929 | 1.05 | 1.12 | 0.0399 | -0.002793
0.9 | 2 | 1.1 | 0.1686 | 0.18546 |  | 0.45 | 0.35165 | 1.12 | 1.05 |  | -1.7372
0.99 | 1.9 | 0.91 | 0.1334 | 0.12139 |  | 2.06 | 1.0343 | 1.2 | 0.76
1.06 | 1.74 | 0.68 | 0.1013 | 0.068884 |  | 0.6 | 0.19625 | 1.37 | 0.6
1.1 | 1.6 | 0.5 | 0.0711 | 0.03555 |  | 0.76 | 0.080089 | 1.69 | 0.45
1.1 | 1.5 | 0.4 | 0.0422 | 0.01688 |  | 1.37 | 0.10693 | 2.06 | 0.13
1.21 | 1.4 | 0.19 | 0.014 | 0.00266 |  | 1.043 | 3.052 |  |  |  | 0.98883
1.4 | 1.21 | -0.19 |  | 3.4954 |  | Average | SUM
1.5 | 1.1 | -0.4
1.6 | 1.1 | -0.5 | W | 0.90713 | if greater than 0.868 then treat as normal
1.74 | 1.06 | -0.68 |  |  | 0.01-level critical point for n = 20
1.9 | 0.99 | -0.91
2 | 0.9 | -1.1
2 | 0.76 | -1.24
2.04 | 0.7 | -1.34
3.1 | 0.4 | -2.7
3.73 | 0.36 | -3.37
0.84195
INDEPENDENT CALIBRATION VERIFICATION
0.1 | 3.1 | 3 | 0.4734 | 1.4202
0.1 | 3 | 2.9 | 0.3211 | 0.93119
0.2 | 2 | 1.8 | 0.2565 | 0.4617
0.2 | 1.7 | 1.5 | 0.2085 | 0.31275
0.4 | 1.5 | 1.1 | 0.1686 | 0.18546
0.5 | 1.4 | 0.9 | 0.1334 | 0.12006
0.9 | 1.3 | 0.4 | 0.1013 | 0.04052
0.9 | 1.2 | 0.3 | 0.0711 | 0.02133
1 | 1.2 | 0.2 | 0.0422 | 0.00844
1 | 1 | 0 | 0.014 | 0
1 | 1 | 0 |  | 3.5016
1.2 | 1 | -0.2
1.2 | 0.9 | -0.3 | W | 0.89858 | if greater than 0.868 then treat as normal
1.3 | 0.9 | -0.4 |  |  | 0.01-level critical point for n = 20
1.4 | 0.5 | -0.9
1.5 | 0.4 | -1.1
1.7 | 0.2 | -1.5
2 | 0.2 | -1.8
3 | 0.1 | -2.9
3.1 | 0.1 | -3
0.84746
LABORATORY CONTROL SAMPLE
0.1 | 24 | 23.9 | 0.4734 | 11.3143
0.3 | 20 | 19.7 | 0.3211 | 6.3257
0.4 | 15 | 14.6 | 0.2565 | 3.7449
0.4 | 13 | 12.6 | 0.2085 | 2.6271
0.5 | 11 | 10.5 | 0.1686 | 1.7703
0.5 | 5 | 4.5 | 0.1334 | 0.6003
0.6 | 4 | 3.4 | 0.1013 | 0.34442
1.5 | 4 | 2.5 | 0.0711 | 0.17775
1.5 | 3 | 1.5 | 0.0422 | 0.0633
1.7 | 2.2 | 0.5 | 0.014 | 0.007
2.2 | 1.7 | -0.5 |  | 26.975
3 | 1.5 | -1.5
4 | 1.5 | -2.5 | W | 0.74271 | if greater than 0.868 then treat as normal
4 | 0.6 | -3.4 |  |  | 0.01-level critical point for n = 20
5 | 0.5 | -4.5
11 | 0.5 | -10.5
13 | 0.4 | -12.6
15 | 0.4 | -14.6
20 | 0.3 | -19.7
24 | 0.1 | -23.9
7.1808
MATRIX INTERFERENCE SAMPLE
-24 | 30 | 54 | 0.4734 | 25.5636
-13 | 20 | 33 | 0.3211 | 10.5963
0.1 | 17 | 16.9 | 0.2565 | 4.3348
0.4 | 12 | 11.6 | 0.2085 | 2.4186
1.4 | 12 | 10.6 | 0.1686 | 1.7872
1.4 | 8 | 6.6 | 0.1334 | 0.88044
1.5 | 5 | 3.5 | 0.1013 | 0.35455
2.7 | 5 | 2.3 | 0.0711 | 0.16353
3.5 | 3.7 | 0.2 | 0.0422 | 0.00844
3.6 | 3.7 | 0.1 | 0.014 | 0.0014
3.7 | 3.6 | -0.1 |  | 46.1089
3.7 | 3.5 | -0.2
5 | 2.7 | -2.3 | W | 0.9007 | if greater than 0.868 then treat as normal
5 | 1.5 | -3.5 |  |  | 0.01-level critical point for n = 20
8 | 1.4 | -6.6
12 | 1.4 | -10.6
12 | 0.4 | -11.6
17 | 0.1 | -16.9
20 | -13 | -33
30 | -24 | -54
11.146

## Trend
ICS | ICV | LCS | MIS
1.1 | 0.5 | 4 | 12
0.76 | 0.1 | 0.5 | 1.4
0.36 | 1 | 1.5 | 8
2.04 | 1.2 | 1.7 | 3.7
0.99 | 0.2 | 0.1 | 12
1.21 | 0.4 | 2.2 | 0.4
1.74 | 1.2 | 0.4 | 3.6
3.73 | 0.9 | 0.3 | 0.1
1.06 | 0.1 | 0.5 | 2.7
3.1 | 1.3 | 15 | 17
2 | 0.9 | 20 | 30
0.7 | 1 | 0.4 | 3.7
0.4 | 2 | 4 | 1.5
0.9 | 0.2 | 0.6 | 5
1.4 | 1 | 1.5 | 1.4
1.9 | 1.4 | 5 | 20
2 | 1.5 | 24 | 3.5
1.5 | 1.7 | 3 | 5
1.6 | 3 | 13 | -24
1.1 | 3.1 | 11 | -13